ExTAG DS 2002/002A

2008-03-19

RECORD OF IECEx / ExTAG DECISIONS

| Standard: IEC 61779-1: 1998-04 Sheet No. 20021002 | Clause: 3.5 | Sheet No. 2002/002A |
| --- | --- | --- |
| Subject: | Key words: | Decision of meeting Revised: 2008 03 19 - During the Denver 2007 Meeting it was decided to revise this DS to take into account changes within the new edition of the Standard. |
| Diffusion Gas Sensing Heads Status of document: adopted | - Diffusion - gas sensing heads - breathing and draining | Date: 2002-10-21 Originator of proposal: AU TC/SC involved: T31 |
| Question: What requirements should be applied for diffusion gas sensing heads as referred to under Clause 3.5 of 1EC61779-1? These sensing heads that contain heated elements, the temperature of which may cause the equipment to exceed the temperature classification. Answer: The applicable requirements of lEC 60079-0:1998 together with the breathing and draining devices requirements of IEC 60079-1:2001 shall be applied. Note: This decision is only intended to address the suitability for use of such equipment in Zone 1. Note : The IEC 61779-1:1998 has been superseded by IEC 60079-29-1 (Ed1.0):2007. In this new standard clause 4.1.2 has now added the following information: “In addition, parts of flammable gas detection apparatus intended for use in hazardous area shall employ materials, and comply with the construction and explosion protection as specified in the other relevant parts of IEC 60079 series.” Therefore for equipment in Zone 0 shall need to comply with IEC 60079-26 while equipment for Zone 1 can comply with IEC 60079-1 or IEC 60079-11. |  |  |