### INTERNATIONAL ELECTROTECHNICAL COMMISSION (IEC) SYSTEM

FOR CERTIFICATION TO STANDARDS RELATING TO EQUIPMENT FORUSE IN EXPLOSIVE ATMOSPHERES (IECEx SYSTEM)

- Circulated to: ExTAG – IECEx Testing and Assessment Group

- TITLE: Compilation of comments on ExTAG/155/CD - Draft Decision Sheet - Method of test for measurement of capacitance of non-metallic enclosures

INTRODUCTION

- This document is a compilation of comments received relating to document ExTAG/155/CD Draft Decision Sheet - Method of test for measurement of capacitance of non-metallic enclosures and is issued for information.

- The document contains responses from the originator (UL), to the comments on ExTAG/155/CD, and, given that these comments are primarily related to minor editorial improvements, with the agreement of the ExTAG Chairman + Secretary, and having taken into account these comments, a revised version of ExTAG/155/CD has been published as ExTAG Decision Sheet 2009/001 - Method of test for measurement of capacitance of non-metallic enclosures.

- The Decision Sheet is now available on the IECEx Web Site

## Michel Brenon

ExTAG Secretary

| Address: IECEx Secretariat 286 Sussex Street Sydney NSW 2000 Australia | Contact Details: Tel: +61 2 8206 6940 Fax: +61 2 8206 6272 E-mail: chris.agius@iecex.com http://www.iecex.com | ExTAG Secretary Mr Michel Brenon/LCIE Tel: +33 1 40 95 5519 Fax: +33 1 40 95 5520 e-mail michel.brenon@lcie.fr |
| --- | --- | --- |

| ExTAG Member | Clause/ Subclause | Paragraph Figure/ Table | Type of comment (General/ Technical/Editorial) | COMMENTS | Proposed change | OBSERVATIONS on each comment submitted |
| --- | --- | --- | --- | --- | --- | --- |
| Electrosuisse (CH) |  |  | General | Electrosuisse have no comments to this Decision Sheet and accept its content. |  | Noted |
| FM Approvals (US) |  |  | General | FM Approvals LLC supports ExTAG/155/CD as drafted. |  | Noted |
| KEMA (NL) |  |  | General | No comments to the Sheet. On the contrary: it covers exactly the problems that we also have seen when doing the test. My compliments to UL for drafting this excellent clarification. |  | Noted |
| NANIO CCVE (RU) |  |  | General | We support this document as Decision Sheet of ExTAG. |  | Noted |
| UL (US) |  |  | General | UL-USA supports this Draft Decision Sheet (ExTAG/155/CD). |  | Noted |
| UL/DEMKO (DK) |  |  | General | UL Demko supports this Draft Decision Sheet (ExTAG/155/CD) |  | Noted |
| FM Approvals (GB) |  |  | Editorial | Modify item 4 to clarify that the value obtained as a result of the initial test should be subtracted from that in item 3 and that include a statement that the value should be positive for the test to be valid. | 4) Compute the difference between the measurements in steps 2 and 3 and record the value. Subtract the value obtained in 3) from that obtained in 2). The value obtained should be positive. | The draft decision sheet was modified to clarify that the value recorded in step 2 (measurement in air) is to be subtracted from the value recorded in step 3 (measurement on test point). In addition, step 4 was revised to clarify that the absolute value of the difference is to be recorded. |
| FM Approvals (GB) |  |  | Technical | By reference to the “positive” test lead the procedure implies that a DC Capacitance measurement system is used. Is the measurement technique critical to this test? Can an AC measurement be used? | Clarify this requirement. | The measurement technique is critical to this test. It has been shown that DC capacitance measurement tools provide more consistent measurements than AC measurement tools. |
| SIMTARS (AU) |  |  | Technical | The information provided in the Answer section of the Decision Sheet is not in line with the requirement in clause 26.15.1 of IEC 60079-0: 2004 and 2007. The standard requires: - that two fully assembled samples of the apparatus be tested - that the capacitance be measured between each exposed metal part of the apparatus. The draft decision sheet suggests testing of two samples of the plastic enclosure and that the capacitance be measured between each exposed metallic part and the earth-ground plate. Based on the above, it is suggested that the text in the draft DS be amended | The test procedure is as follows: Two fully assembled samples of the apparatus samples of the polymeric enclosure of the device…….…. to 50%.The sample ….necessitates it ).If the sample ……held by hand. .. Using the a battery-powered capacitance meter, the capacitance between each exposed metallic part on the apparatus and the earth-grounded metal plate is to be measured in the range 0 pF to 200 pF with an accuracy of ±5% and connection leads as short as possible but, in any case, less than 1 m. If …metallic test points are not easily accessible to the meter leads, a screw may be inserted to expose the test point. | Two fully assembled samples of the apparatus must be used for the test. The draft decision sheet was modified to clarify this point. The capacitance is to be measured between each exposed metallic part and the earth-ground plate, not between each exposed metallic part. Electrostatic discharges usually occur between a charged object and an earthed object. Thus, the capacitance between exposed metallic points and earth-ground is used to measure the equipment’s ability to store electrostatic charge. Battery-powered capacitance meters are earth-free and provide consistent measurements. |