- INTERNATIONAL ELECTROTECHNICAL COMMISSION (IEC) SYSTEM FOR CERTIFICATION TO STANDARDS RELATING TO EQUIPMENT FOR USE IN EXPLOSIVE ATMOSPHERES (IECEx SYSTEM)

TITLE: Draft ExTAG Decision Sheet – The resistance to impact test according to IEC 60079-0

- Circulated to: ExTAG – IECEx Testing and Assessment Group

INTRODUCTION

This proposed Draft ExTAG Decision Sheet ExTAG/477/CD –The resistance to impact test according to IEC 60079-0, has been prepared by PCEC China.

While initially discussed during the 2017 ExTAG Washington Meeting the decision was taken during that meeting to recirculate the document, without change, for a six week comment period.

Please complete the Comments Table, separate document, and return to the Secretariat

by –

2017 12 11 to

Christine Kane

Julien Gauthier

Julien Gauthier

ExTAG Secretary

| Address: IECEx Secretariat Level 33 Australia Square 264 George Street Sydney NSW 2000 Australia Web: www.iecex.com | ExTAG Secretary Mr Julien Gauthier LCIE S.A. 33 Avenue du General Leclerc 92260 Fontenay-aux-Roses FRANCE Tel: +33 1 40 95 55 26 Fax: +33 1 40 95 89 37 Email : julien.gauthier@fr.bureauveritas.com |
| --- | --- |

COLLECTION OF IECEx / ExTAG DECISION

| Standard: IEC 60079-0:2011 Edition 6.0 | Clause: 26.4.2 26.7.2 |  |
| --- | --- | --- |
| Subject: The resistance to impact test according to IEC 60079-0 Status: Draft | Key words: - impact | Date: 10 August 2017 Originator of proposal: PCEC TC/SC involved: TC 31 (MT 60079-1) |
| Background: In IEC 60079-0:2011 Edition 6.0 Clause 26.4.2, Resistance to impact test should be carried out on Light-transmitting parts. Glass is a kind of common used Light-transmitting part which is an essential test object. As a kind of non-metallic material, glass do not subject to thermal endurance test. But tests shown the lower temperature will affect the mechanical strength of the glass. The temperature requirement for resistance to impact test in Clause 26.4.2 are as follows: The test shall be carried out at an ambient temperature of (20 ± 5) °C, except where the material data shows it to have a reduction in resistance to impact at lower temperatures within the specified ambient range. In this case, the test shall be performed at the lower test temperature, in accordance with 26.7.2. Question: Is it necessary to carry out resistance to impact test of glass at lower temperatures according to clause 26.7.2. Are there any requirement for data and information about glass which should be provided by manufacture. Answer: Yes, in the test of resistance to impact for glass, the samples must be cooled down to the lower test temperature which is specified in clause 26.7.2. Also, data sheet for glass which includes mechanical character should be provided by manufacture. |  |  |