COLLECTION OF IECEx / ExTAG DECISIONS

| Standard: IEC 60079-1 Sixth Edition 2007 | Clauses: C.3.1 |  |
| --- | --- | --- |
| Subject: Sealing test of cable glands Status of document: Approved | Key words: - Cable glands - Sealing test - Thermal endurance | # Date: 2010.10 14 Originator of proposal: D R Stubbings/Sira TC/SC involved: |
| Introduction Clause C.3.1 of IEC 60079-1:2007 requests that cable glands be assembled on to either mandrels or cable (as applicable to their design) in accordance with the manufacturer’s instructions and then subjected to the appropriate thermal conditioning parameters as applicable in accordance with IEC 60079-0. The sealing test is conducted after thermal endurance. Question: Is it permissible to retighten a cable gland in accordance with the manufacturer’s instructions prior to carrying out the sealing test? Answer: Yes: - The test requirement of clause C.3.1 calls for thermal endurance tests on an assembled cable gland - Clause C.3.1.1 follows clause C3.1, this calls for the gland to be tightened to the manufacturer’s specified torque, prior to the sealing test being carried out. |  |  |