INTERNATIONAL ELECTROTECHNICAL COMMISSION SYSTEM FOR
CERTIFICATION TO STANDARDS RELATING TO EQUIPENT FOR USE
IN EXPLOSIVE ATMOSPHERES (IECEx SYSTEM)

- Title: Mid-term and Scope Extension Assessment Report for Intertek Testing Services NA, Inc., Plano, Texas an Accepted ExTL within the IECEx System, Equipment Scheme 02

- To: Members of the IECEx Management Committee, ExMC

Introduction

During the Mid-term assessment of Intertek Testing Services NA, Inc an Accepted Ex Test Laboratory (ExTL), the opportunity was taken to also conduct as assessment for a scope extension application to include the following Standards within their scope.

| 60079-5 | Explosive atmospheres - Part 5: Equipment protection by powder filling 'q' |
| --- | --- |
| 60079-6 | Explosive atmospheres - Part 6: Equipment protection by oil immersion 'o' |
| 60079-18 | Explosive atmospheres - Part 18: Equipment protection by encapsulation "m" |
| 60079-25 | Explosive atmospheres - Part 25: Intrinsically safe systems |
| 60079-26 | Explosive atmospheres - Part 26: Equipment with equipment protection level (EPL) Ga |
| 60079-28 | Explosive atmospheres - Part 28: Protection of equipment and transmission systems using optical radiation (excluding ignition tests to 5.2.4) |
| 60079-31 | Explosive atmospheres - Part 31: Equipment dust ignition protection by enclosure "t" |

This report details the assessment findings of this scope extension with the IECEx Assessment Team recommending the acceptance of the above scope extension. It is now hereby submitted for voting by ExMC.

Please consider the report and return the completed voting form, separate Word document, to the Secretariat by 2015 09 29

Your speedy response to the voting process will be very much appreciated.

Chris Agius

IECEx Secretary

| Address: Level 33, Australia Square 264 George Street Sydney NSW 2000 Australia | Contact Details: Tel: +61 2 46 28 4690 Fax: +61 2 46 27 5285 e-mail:chris.agius@iecex.com http://www.iecex.com |
| --- | --- |

IEC System for certification to standards relating to equipment for use in Explosive Atmospheres (IECEx System)

IECEx Assessment Report Form

IECEx Assessment Report Form for use by IECEx Assessment Teams to report Assessments conducted according to the IECEx Assessment Procedures of

- Operational Document IECEx OD 003-2 for the Certified Equipment Scheme

- Operational Document IECEx OD 016 for the Certified Service Facility Scheme

- Operational Document IECEx OD 022 for the IECEx Conformity Mark Licensing System

IECEx ExTL assessment report for Intertek Testing Services NA, Inc., Plano, Texas

INTERNATIONAL
ELECTROTECHNICAL
COMMISSION

CONTENTS

1	Assessment information	3

1.1	Type of Body covered by this assessment:	3

1.2	Type of assessment:	3

1.3	Details of body	3

1.3.1	Country	3

1.3.2	Name of body	3

1.3.3	Name and title of nominated principal contact	3

1.4	Assessment information	3

1.4.1	Members of the assessment team	3

1.4.2	Place(s) of assessment	3

1.4.3	Assessment date(s)	3

1.5	Scope	3

1.5.1	ExTL scope for equipment certification scheme	3

1.5.2	ExCB scope	4

1.6	Legal entity of body	4

1.7	Financial support	4

1.8	History	4

1.9	Quality manual	4

1.9.1	Procedures	5

1.9.2	Work instructions	5

1.9.3	Records (including test records where relevant)	5

1.9.4	Document change control	5

1.10	Confidentiality	5

1.11	Publications (Hard cover and Electronic)	5

1.12	Recognition and agreements	5

1.13	Internal audit and periodic management review	5

1.14	Contracting, subcontracting, use of other labs and use of other locations	5

1.15	Training and competence	5

1.16	Complaints and appeals (including appeals to IECEx)	5

1.17	Special facts to be noted	5

1.17.1	Supporting documentation	5

1.17.2	Tests Witnessed	5

1.18	Issues raised	6

1.19	Recommendations	6

2	ExTL for IECEx Certified Equipment Scheme	7

2.1	Assessment references	7

2.2	Candidate ExTL persons interviewed	7

2.3	Associated ExCB(s)	7

2.4	Organisation	7

2.5	Organizational structure	7

2.6	Resources	7

2.7	Test reports issued	7

2.8	National accreditation	8

2.9	Calibration	8

2.10	Comments (including issues found during assessment)	8

Annex A A2LA Accreditation	9

# Assessment information

## Type of Body covered by this assessment:

| ExCB for IECEx Certified Equipment Scheme |  |
| --- | --- |
| ExTL for IECEx Certified Equipment Scheme |  |
| ExCB for IECEx Certified Service Facilities Scheme |  |
| ExCB for IECEx Conformity Mark Licensing System |  |

NOTE 1	ExCB - IECEx Certification Body

NOTE 2	ExTL - IECEx Testing Laboratory

## Type of assessment:

| Pre-assessment for candidate body |  |
| --- | --- |
| Initial assessment for candidate body |  |
| Surveillance |  |
| Re-assessment |  |
| Scope extension |  |

## Details of body

### Country

US

### Name of body

Intertek Testing Services NA, Inc., Plano, Texas.

### Name and title of nominated principal contact

| Name | Title | E-mail address |
| --- | --- | --- |
| Ryan Parks | Engineering Manager | ryan.parks@intertek.com |

## Assessment information

### Members of the assessment team

| Name | Role (modify as necessary) |
| --- | --- |
| Ron Webb | IECEx Lead Assessor |
| Lucy Lu | Expert Assessor |

### Place(s) of assessment

1809 10th Street, Suite 400,Plano, TX 75074, USA

### Assessment date(s)

19-20 May 2015

## Scope

### ExTL scope for equipment certification scheme

Items below greyed out are already in scope.

| Number | Title |
| --- | --- |
| 60079-0 Edition 6 | Explosive atmospheres - Part 0: Equipment - General requirements |
| 60079-1 Edition 7 | Explosive atmospheres - Part 1: Equipment protection by flameproof enclosures “d” |
| 60079-2 Edition 6 | Explosive atmospheres - Part 2: Equipment protection by pressurized enclosure “p” |
| 60079-7 Edition 4 | Explosive atmospheres - Part 7: Equipment protection by increased safety "e" |
| 60079-11 Edition 6 | Explosive atmospheres - Part 11: Equipment protection by intrinsic safety “i” |
| 60079-15 Edition 4 | Explosive atmospheres - Part 15: Equipment protection by type of protection “n” |

Standards which scope extension is requested

| Number | Title |
| --- | --- |
| 60079-5 Edition 4 | Explosive atmospheres - Part 5: Equipment protection by powder filling 'q' |
| 60079-6 Edition 3 | Explosive atmospheres - Part 6: Equipment protection by oil immersion 'o' |
| 60079-18 Edition 4 | Explosive atmospheres - Part 18: Equipment protection by encapsulation "m" |
| 60079-25 Edition 2 | Explosive atmospheres - Part 25: Intrinsically safe systems |
| 60079-26 Edition 3 | Explosive atmospheres - Part 26: Equipment with equipment protection level (EPL) Ga |
| 60079-28 Edition 1 | Explosive atmospheres - Part 28: Protection of equipment and transmission systems using optical radiation (excluding ignition tests to 5.2.4) |
| 60079-31 Edition 2 | Explosive atmospheres - Part 31: Equipment dust ignition protection by enclosure "t" |

NOTE 1: Unless otherwise indicated, earlier editions of standards (even if with a different number) are considered to be covered in the above scope for the purposes of the assessment.

### ExCB scope

Intertek Testing Services NA, Inc., Cortland, US covers all the standards listed in 1.5.1 with the exception of IEC 60079-28 which is covered by Intertek Testing and Certification Ltd, Leatherhead, UK.

## Legal entity of body

Not relevant as this is a scope extension assessment.  This was covered in the initial assessment reported in ExMC_850_DV.

## Financial support

Not relevant as this is a scope extension assessment.   This was covered in the initial assessment reported in ExMC_850_DV.

## History

Not relevant as this is a scope extension assessment.  This was covered in the initial assessment reported in ExMC_850_DV.

Documentation

## Quality manual

Not relevant as this is a scope extension assessment.  This was covered in the initial assessment reported in ExMC_850_DV.

### Procedures

Relevant procedures for the new standards were reviewed during the assessment visit and found to meet the requirements of IECEx

### Work instructions

Relevant work instructions for the new standards were reviewed during the assessment visit and found to meet the requirements of IECEx

### Records (including test records where relevant)

Not relevant as this is a scope extension assessment. This was covered in the initial assessment reported in ExMC_850_DV.

### Document change control

Not relevant as this is a scope extension assessment..  This was covered in the initial assessment reported in ExMC_850_DV.

## Confidentiality

Not relevant as this is a scope extension assessment. This was covered in the initial assessment reported in ExMC_850_DV.

## Publications (Hard cover and Electronic)

Not relevant as this is a scope extension assessment.  This was covered in the initial assessment reported in ExMC_850_DV.

## Recognition and agreements

Not relevant as this is a scope extension assessment.  This was covered in the initial assessment reported in ExMC_850_DV.

## Internal audit and periodic management review

Not relevant as this is a scope extension assessment. This was covered in the initial assessment reported in ExMC_850_DV. However it has been expanded to cover new Standards

## Contracting, subcontracting, use of other labs and use of other locations

No subcontracting for the standards in scope.

## Training and competence

There is a competency matrix for the IECEx Scheme.    Interviews were undertaken with staff from the testing section to ensure that they had the required level of understanding.  These are included in a TCD document which is held by the Secretariat. Training records for each staff member are held on a SharePoint system and are very comprehensive.

## Complaints and appeals (including appeals to IECEx)

Not relevant as this is a scope extension assessment. This was covered in the initial assessment reported in ExMC_850_DV.

## Special facts to be noted

### Supporting documentation

- Information concerning the site visit is included in a Site Assessment Report from this visit which is held by the Secretariat.

### Tests Witnessed

Tests witnessed in connection with IEC 60079-18, IEC 60079-28 and IEC 60079-31 are as follows:

Dielectric strength test on encapsulant material – 60079-18 clause 8.1.2

Dust test for Ex ta – IEC 60079-31 clause 6.1.1.4

Radiance measurement for op is equipment IEC 60079-28 clause 5.2.2

## Issues raised

There were no formal issues raised.  A few minor details concerning local operating procedures were cleared during the assessment.

## Recommendations

Based on the assessment performed on 19-20 May 2015, Intertek Testing Services NA, Inc., Plano, Texas is recommended for scope extension for the following standards.

| Number | Title |
| --- | --- |
| 60079-5 Edition 4 | Explosive atmospheres - Part 5: Equipment protection by powder filling 'q' |
| 60079-6 Edition 3 | Explosive atmospheres - Part 6: Equipment protection by oil immersion 'o' |
| 60079-18 Edition 4 | Explosive atmospheres - Part 18: Equipment protection by encapsulation "m" |
| 60079-25 Edition 2 | Explosive atmospheres - Part 25: Intrinsically safe systems |
| 60079-26 Edition 3 | Explosive atmospheres - Part 26: Equipment with equipment protection level (EPL) Ga |
| 60079-28 Edition 1 | Explosive atmospheres - Part 28: Protection of equipment and transmission systems using optical radiation (excluding ignition tests to 5.2.4) |
| 60079-31 Edition 2 | Explosive atmospheres - Part 31: Equipment dust ignition protection by enclosure "t" |

This is according to the scope of the standards listed in this document.

| Ron Webb | Lucy Lu |
| --- | --- |
| Lead Assessor | Expert Assessor |

Date:  20 May 2015

# ExTL for IECEx Certified Equipment Scheme

## Assessment references

- IECEx02 IECEx Certified Equipment Scheme covering equipment for use in explosive atmospheres – Rules of Procedure

- IECEx OD003-2 Assessment, surveillance assessment and re-assessment of ExCBs and ExTLs operating in the IECEx 02, IECEx Certified Equipment Scheme

- IECEx OD009 Issuing of CoCs, ExTRs and QARs

- ISO/IEC 17025:2005 Edition 2, General requirements for the competence of testing and calibration laboratories

- IECEx Document OD17 Drawing and documentation guidance

- IECEx Technical Capability Documents (TCDs)

- ExTAG decision sheets (DSs)

NOTE	The latest editions of the above documents were applied.

## Candidate ExTL persons interviewed

Several members of the staff were interviewed and details are recorded in the Site Assessment Report

## Associated ExCB(s)

Associated ExCB for all Standards excluding 60079-28 is Intertek Testing Services NA, Inc. Cortland, US

For 60079-28 the associated ExCB is Intertek Testing and Certification Ltd, Leatherhead, UK.

## Organisation

Not relevant as this is a scope extension. 	This was covered in the initial assessment reported in ExMC_850_DV.

## Organizational structure

Not relevant as this is a scope extension assessment. This was covered in the initial assessment reported in ExMC_850_DV.

## Resources

The operation is well resourced with appropriate facilities, test equipment, procedures and experienced staff to cover the scope extension

## Test reports issued

Number of test reports issued for the preceding two years for each type of protection for the scope extension

| Standard numbers | Type of protection or other identifying information | Number of issued reports (for last 2 years) |
| --- | --- | --- |
| 60079-0 | General requirements | Included as appropriate |
| 60079-5 | Ex q | 0 |
| 60079-6 | Ex o | 2 |
| 60079-18 | Ex m | 0 |
| 60079-25 | Intrinsically safe systems | 2 |
| 60079-26 | Ga | 6 |
| 60079-28 | Ex op | 1 |
| 60079-31 | Ex t | 2 |

## National accreditation

Intertek Testing Services NA, Inc. Plano holds accreditation from A2LA to IEC 17025 which covers the following standards:
IEC 60079-0, IEC 60079-1, IEC 60079-2, IEC 60079-7, IEC 60079-11 and IEC 60079-15.

Until this scope is extended to cover the standards listed in 1.5.1, annual surveillance will  be required.

## Calibration

Calibration details for equipment used for the scope extension was viewed and found to meet the requirements of the IECEx.

## Comments (including issues found during assessment)

None.

- A2LA Accreditation